Protocol for Radiographs

- Apron, thyroid shield, gloves, and dosimeter monitoring device must be worn for every x-ray, every time. Fingers shown in an x-ray will be considered a disciplinary matter.

- Monitoring device must be worn at the thyroid level on the collar outside the apron.

- Failure to limit or restrict the X-ray beam only to the area of clinical interest represents one of the most frequent causes of violating the collimation requirement. X-rays that extend beyond the area of clinical interest serve no useful function, increase scatter, and must be eliminated by careful collimation.

- The operator shall not stand in the beam, but be well away from the tube and animal during X-raying.

- Proper storage of aprons and gloves prolong their life and effectiveness. Aprons should be hung without creases to prevent cracking.

- Gloves and aprons are designed to protect the wearer from scatter radiation only. They do not reduce the primary X-ray beam enough to provide sufficient protection.

- Use high kVp techniques that are appropriate for the body part being X-rayed, permitting the operator to lower the mAs settings, therefore decreasing radiation levels. Kilovoltage determines the penetrating ability (quality) of the X-ray beam whereas mAs determines the amount (quantity) of radiation.

- Plan radiographic procedures carefully and avoid unnecessary retakes. Fractious animals should not even have radiographs attempted on them without sedation.

- Retakes represent one of the biggest causes of excessive and unnecessary radiation exposure to the patient and personnel.

- The body part should be perpendicular to the central main X-ray beam and parallel to the film.

- Ensure the focal-film distance in correct (40 inches)

- Appropriate physical Left and Right markers must be used for every x-ray.

- The more senior employee taking x-rays must enter their name as “Tech” in acquisition software.

| Received & Understood by: |  | Signature: |  | Date: |  |
| --- | --- | --- | --- | --- | --- |